KENNETH S. FINK, MD, MGA, MPH

DIRECTOR OF HEALTH

KA LUNA HOʻOKELE

#### JOSH GREEN, M.D.

GOVERNOR OF HAWAIʻI

KE KIAʻĀINA O KA MOKUʻĀINA ʻO HAWAIʻI

In reply, please refer to:

File:SLD/ADMIN

STATE OF HAWAIʻI

## DEPARTMENT OF HEALTH

## State Laboratories Division (SLD)

### 2725 Waimano Home Road

### Pearl City, HI 96782

February 19, 2026

The Department of Health, State Laboratories Division (SLD), Laboratory Preparedness and Response Branch (LPRB) is requesting a quote for a temperature and humidity monitoring system at the Kamaule’ule’ Building, 2725 Waimano Home Road, Pearl City, HI 96782

I.  Bidder Requirements

The Bidder shall:

A.  shall be both the manufacturer and installer for the system.

B.  shall become familiar with all the job requirements prior to the submission of a quote.

C.  shall provide documentation demonstrating previous experience with the 	installing and maintaining the laboratory temperature monitor system in laboratories of	similar or greater size and complexity to SLD and shall provide at least three 	references from existing customers within the United States upon request of SLD.  A 	minimum of one (1) of the references shall be from an ISO/IEC 17025 accredited 	laboratory using the system as part of their quality management system.

D.  shall ensure that it is current with all payments and registration fees and similar 	financial obligations owed to the STATE at the time of award and time of final payment;

E.  shall include adequate and sufficient documentation with bid submission for the State 	to verify the bidder is able to meet the requirements listed in this Solicitation.

F.  shall be able to work within the IT requirements of SLD and provide documentation 	of any internet access/computer requirements prior to award.  If bidder desires to have 	any IT 	requirements reviewed prior to submitting a bid please submit those requirements 	to the contacts listed at the end of the solicitation no later than February 28, 2026.

G.  shall not subcontract any of the installation, testing, training, warranty, and 	extended warranty for the system.

H.  shall submit any questions or request for a site visit to the contacts listed in section IV 	no later than 2:00 PM on February 28, 2026,

I.   shall provide a certificate of insurance (COI) prior to the start of any onsite 	installation work.  Please see attached template for insurance and COI requirements.

J.  shall agree to the standard State of Hawaii General Terms and Conditions no 	modifications shall be allowed.

II.  System Requirements

The System shall:

A.  be provided in new condition, which includes but is not limited to the latest version 	of hardware and software for the system.

B.  monitor forty-nine (49) pieces of  equipment and twenty-one (21) rooms listed in the 	attached Equipment List and be capable of expansion up to 500+ pieces of equipment or 	locations.  Monitoring shall include temperature and the open/closed status of the listed 	freezers, refrigerators, and incubators and the temperature and humidity of the listed 	rooms.

C.  be cloud-based allowing for secured remote access to the monitoring system data at 	any time.

D.  not require modification (drilling or other destructive methods) to properly install 	sensors into the monitored equipment and not require installation of wired connections 	from individual sensors back to a server or aggregation point.

E.  be designed specifically for real-time environmental parameter sensing and 	monitoring;

F.  configurable to allow for comprehensive alarm notifications, event logging, and 	investigative reporting capabilities; Alarms shall at minimum be audible, email, and 	phone.

G.  installation shall be primarily wireless, and scalability based on the STATE 	monitoring needs. A system that has monitors and probes attached to individual units 	which then communicate wirelessly with a limited number of hard-wired access points 	located in the communication or electrical rooms would be acceptable.

H.  be able to buffer data for a minimum of seven (7) days in the event of a power loss or 	communication loss.

I.  shall be accessible and usable on fully Windows-based systems and provide security 	options that include monitoring system-specific passwords.

J.  shall maintain all records in the cloud for a minimum of five  (4) years.

III.  Scope of work

The Contractor shall:

A.  provide all necessary materials, labor, equipment and services to deliver, install, test, 	and providing ongoing calibration and maintenance of the System.  Materials shall 	include but are not limited to all required hardware, software, wiring, probes and other 	incidentals required to complete installation and connectivity.  Contractor shall be solely 	responsible for any travel and lodging costs.

B.  be responsible for verifying field conditions, final service connections and locations.

C.  Install, test and validate the system on-site using only certified technician(s) directly 	employed by the manufacturer.

D.  be responsible for protecting laboratory equipment and employee work areas, 	cleaning the work area daily, that includes removal of trash, unused supplies, debris and 	visible dust created by the work performed, and for the haul-away and proper disposal of 	all discarded materials associated with the job. Contractor shall be responsible for the 	proper and lawful disposal of all other incidental debris produced from this job, per 	typical construction waste management practice.

E.  responsible for meeting all required Federal, State, and County laws including all 	Occupational Safety and Health Administration (OSHA) and Hawaii Occupational Safety 	and Health Division (HIOSH) requirements pertaining to personnel safety for the work 	involved.

F.  maintain effective communication with the STATE and cooperate in the resolution of 	problems, suspected problems, or potential problems.

G.  conduct the testing and validation of the System according to company’s established 	procedures and shall provide a copy of the procedure to the STATE upon request.  The 	testing of the System shall at a minimum include verification of sensor functionality, 	calibration, alarm functionality and communication functionality.

H.  document all test and validation results and shall provide a copy of the report to the 	STATE upon completion.  Test, validation and calibration reports shall verify 	conformance to specified manufacturer’s performance requirements have been met.

I.  support the SLD-LPRB's compliance with ISO/IEC 17025 requirements, including but 	not limited to traceable calibration of sensors, documentation of measurement of 	uncertainty, and audit trail capability.

J.  calibrate temperature and humidity sensors on-site utilizing National Institute of 	Standards and Technology (NIST) traceable equipment and shall provide a copy of the 	NIST certificate to the STATE.

K.  provide all training necessary to monitor and manage the system to STATE staff

post-installation on the System. Training shall:

1. be scheduled with SLD within two (2) weeks of the installation and validation 			of the system.

2.  be documented by the Installer, including names of personnel, training agenda 			and date of training.

3.   include hands-on training with the System which shall include 	but is not 			limited to demonstrations on how to set up alarms, respond to alarms, and run 			reports.

4.  include one (1) hard copy and an electronic copy of all operating and 				maintenance manual for each item installed and for the system as a whole.

L.  provide drawings showing the location of all equipment as installed within the 	building.

M.  provide one (1) year warranty of all parts and installation work.  Warranty shall:

1.  include all associated travel and lodging costs with onsite maintenance of the 			system.  Warranty shall include any battery replacements.

2.  One (1) calibration of each probe shall be included as part of the Warranty and 			be conducted within one month prior to the end of the warranty period.  				Calibration shall be done onsite using NIST traceable equipment.

3.  documentation of all warranty and calibration shall be provided to SLD.

N.  provide up to five (5) years of extended warranty/service at a yearly price that shall 	be firm for a minimum of five (5) years.  Extended warranty shall:

1.  include all labor and parts required to keep the system in regular working order 		and include all associated travel and lodging costs associated with onsite 				maintenance of the system.  Extended warranty shall include any battery 				replacements.

2.  Yearly calibration of each probe shall be included as part of the extended 			warranty and be conducted during the month of the previous year’s calibration.  			Calibration shall be done onsite using NIST traceable equipment.

3.  documentation of all extended warranty work and calibration shall be provided 		to SLD.

O.  complete delivery, installation, and training by May 31, 2026

IV.  Bid Submission

Please submit your price quotation by 12 PM (Hawaii Time), March 6, 2026, on HIePRO. Award will be made to the lowest responsive, responsible offer. The offeror must be able to produce a Certificate of Vendor Compliance at the time of award and for final payment.

For information on the Certificate of vendor Compliance, please refer to the SPO Website at www.spo.hawaii.gov

When entering your bid on HIePRO.  Enter the total cost of the system, installation, training, and initial warranty inclusive of taxes and fees for line item 1.  Then enter the yearly cost of extended warranty inclusive of any taxes and fees for line item 2.  The system should multiply the yearly cost by five (5) for the total amount for the line item.  Attach documentation confirming the ability of the system to meet the requirements of this solicitation, any IT requirements, and the required references.

While not required it is highly encouraged to attach a quote that confirms the accuracy of the numbers entered on HIePRO.

The prices shall include all applicable taxes and fees. HIePRO may charge transaction fees that the awarded bidder shall be responsible for paying.

If you have any questions, require more information, would like to request a site visit, or would like to submit documentation ahead of time please contact Mr. Stephen Schanzenbach at (e-mail: stephen.schanzenbach@doh.hawaii.gov or phone at 808-453-6651).

V.  Method of Award

Award will be to the lowest responsive, responsible offeror.  The lowest offer will be determined by the cost of the system + the number of years of extended warranty selected by the Procurement Officer based on the availability of funding.  SLD intends to make a single award to the offeror that can provide the most years of extended warranty plus the cost of the system up to a maximum of five (5) years of extended warranty while being under the maximum budget for this project.  The maximum budget for this project is $99,999.99.

VI.  Other Information

The Hawai‘i Department of Health does not discriminate on the basis of race, color, sex, national origin, age, or disability, or any other class as protected under applicable federal or state law, in administration of its programs, or activities, and, the Department of Health does not intimidate or retaliate against any individual or group because they have exercised their rights to participate in actions protected, or oppose action prohibited, by 40 C.F.R. Parts 5 and 7, or for the purpose of interfering with such rights.

If you have any questions about this notice or any of the Department’s non-discrimination programs, policies, or procedures, you may contact:

Valerie Kato

Acting Non-Discrimination Coordinator

Hawai‘i Department of Health

1250 Punchbowl Street, HI 96813,

(808) 586-4400

doh.nondiscrimination@doh.hawaii.gov

If you believe that you have been discriminated against with respect to a Department of Health program or activity, you may contact the Non-Discrimination Coordinator identified above.

To request language or accessibility for this document, please contact: the HDOH Non-Discrimination Coordinator, located at 1250 Punchbowl Street, Honolulu, HI 96813 (Phone: (808) 586-4400 or email: doh.nondiscrimination@doh.hawaii.gov).

Please allow sufficient time for HDOH to meet accommodation requests.