KENNETH S. FINK, MD, MGA, MPH

DIRECTOR OF HEALTH

KA LUNA HOʻOKELE

#### JOSH GREEN, M.D.

GOVERNOR OF HAWAIʻI

KE KIAʻĀINA O KA MOKUʻĀINA ʻO HAWAIʻI

In reply, please refer to:

File:

STATE OF HAWAIʻI

## DEPARTMENT OF HEALTH

## State Laboratories Division (SLD)

### 2725 Waimano Home Road

### Pearl City, HI 96782

April 20, 2026

REQUEST FOR PRICE QUOTATIONS FOR CERTIFICATION OF FUME HOODS AND BIOSAFETY CABINETS

The Hawaii State Laboratories Division is requesting price quotations for the testing, certification, adjustment using the National Sanitation Foundation (NSF) International Standard 49 for biological safety cabinets (BSCs) and ANSI  Z9.5 for fume hoods and trace metal hoods located at the State Laboratories Facility at 2725 Waimano Home Road, Pearl City, HI 96782.

Included in this proposal are certification of 40 BSCs, 9 Laminar Flow benches, and 67 fume hoods (116 pieces of equipment total).

Besides testing and certification, offeror shall be responsible for decontamination of BSCs if it is necessary for certification.

The general breakdown is as follows:

| Line No. | Equipment | Qty |
| --- | --- | --- |
| 1 | Baker BSC II | 6 |
| 2 | Fume Hoods | 67 |
| 3 | Labconco BSC II | 30 |
| 4 | Laminar Flow Clean Bench | 9 |
| 5 | Nuaire BSC II | 4 |

Scope of Services

- Offeror shall preform the required tests to certify the equipment listed in the attached equipment list.

- Offeror shall provide a final report (Test and Certification Report) and specific/individual Performance Evaluation Report for each BSC/Fume Hood following recertification.

- Each report shall clearly indicate on the first page of the report the operational status the equipment was left in and its suitability to be used for laboratory work.  Each report shall clearly indicate each test performed and the results of those tests.  For equipment which did not pass one or more of the tests the report shall indicate what repairs are required.

- Equipment Reports shall include the results of each test done to certify the equipment.  For BSCs this shall include down flow velocity, inflow velocity, airflow pattern tests, HEPA/ULPA filter leak test, sash alarm, exhaust airflow alarms and interlocks tests.  For Fume hoods, this shall include face velocity measurement and smoke visualization tests.  In the case of all equipment, certifier shall also include the results any other tests performed but were not listed above that would be required to certify the equipment to the applicable standard.

- Each report shall include a list of equipment (including model and serial numbers) used during the certification of the BSC or fume hood.  A copy of the current certification certificates for all equipment to be used in testing shall be provided to the State prior to the start of work upon request.

- Offeror shall immediately notify the Building Manager verbally of any BSCs, Fume Hoods, or laminar flow hoods that fail recertification.  A written summery of the any failed equipment and recommended repairs needed to bring units back to working order within a week of the failed test.

- If a BSC/Fume Hood passes recertification, Offeror shall place a signed sticker label on each BSC/Fume Hood indicating the date of recertification, Serial Number of the BSC/hood, the signature of the person who performed the evaluation and any other relevant information.

- If a BSC/Fume Hood fails recertification, Offeror shall place a sticker that is visually distinct from the sticker used on BSC/Fume Hoods that pass recertification that clearly states the BSC/Fume Hood is not to be used.

- Offeror shall be responsible for safe disposal for any waste generated as a result of certification or decontamination.

Time of Performance - Offeror shall perform the certification work starting no sooner than June 1, 2026 and complete all certifications before June 30, 2026.  Work hours shall be between 8 AM and 4 PM on Weekdays (excluding June 11 which is a State Holiday).  Equipment reports shall be provided to the STATE for initial review by July 20, 2026.

Qualifications to Bid - Offeror shall be an NSF Accredited Biosafety Cabinet Field Certifier or have an NSF Accredited Biosafety Cabinet Field Certifier on staff.  All BSC certification work shall be done by individuals who are current NSF Accredited Biosafety Cabinet Field Certifier. Please list the name(s) and the certificate number(s) of those who will be onsite on the worksheet.

Quote Requirements - Offeror shall provide a quotation that includes cost/unit (inclusive of tax and any fees).  Offeror shall complete attached worksheet and upload the completed work sheet with their offer on the electronic bid system.  Total price (inclusive of taxes) on worksheet must equal the total entered into the electronic bid system.  SLD reserves the right to add additional units for certification at the unit prices listed on the worksheet during the service period (June 1, 2026 – June 30, 2026).

Method of award - Award shall be made to the Offeror that provides the lowest responsive, responsible offer.  Offeror must be compliant with all Federal and State laws and be compliant with the State of Hawaii on Hawaii Compliance Express (HCE) to be eligible for award and at the time of final payment.

Other Information

The Hawai‘i Department of Health does not discriminate on the basis of race, color, sex, national origin, age, or disability, or any other class as protected under applicable federal or state law, in administration of its programs, or activities, and, the Department of Health does not intimidate or retaliate against any individual or group because they have exercised their rights to participate in actions protected, or oppose action prohibited, by 40 C.F.R. Parts 5 and 7, or for the purpose of interfering with such rights.

If you have any questions about this notice or any of the Department’s non-discrimination programs, policies, or procedures, you may contact:

Valerie Kato

Acting Non-Discrimination Coordinator

Hawai‘i Department of Health

1250 Punchbowl Street, HI 96813,

(808) 586-4400

doh.nondiscrimination@doh.hawaii.gov

If you believe that you have been discriminated against with respect to a Department of Health program or activity, you may contact the Non-Discrimination Coordinator identified above.

To request language or accessibility for this document, please contact: the HDOH Non-Discrimination Coordinator, located at 1250 Punchbowl Street, Honolulu, HI 96813 (Phone: (808) 586-4400 or email: doh.nondiscrimination@doh.hawaii.gov).

Please allow sufficient time for HDOH to meet accommodation requests.