Lab Equipment: Autoclave Specification Requirements for HiePRO purchase

The Natural Energy Laboratory of Hawaii Authority (NELHA), an attached agency of the Department of Business, Economic Development and Tourism is soliciting bids for the Water Quality Laboratory to purchase an autoclave.  The unit shall allow for a complete sterilizing cycle that is fully automatic, time controlled with built in temperature control providing 121oC, have easy access operating controls, pressure is automatically exhausted and the power supply is shut off at the end of the cycle. The preferred equipment size is 85L (3 cubic feet).

A complete autoclave / sterilizer must have the following features:

- Thermometer 180 to 300 °F (82 to 149 °C)

- 0 to 30 lbs. steam gauge

- 0 to 60 min timer

- Thermostatic steam trap

- Signal light

- Flat perforated steam baffle

- Removable pan supports

- Adjustable 4 to 6" bullet feet

- Slow and fast exhaust options

- Manual fill of chamber

- Safety valve

- Low-water cut-off

- Automatic temperature control

- Perforated steam baffle

The following options are not necessary:

- Recording thermometer

- Printer capability

- USB port for recording settings

A product that would meet these requirements is manufactured by Market Forge, catalog #STM-E; similar products will be considered.